Real-Time Acceleration for Medical Image Processing

Project documentation, version 2.0 (after the MathWorks review)

Project 259, MathWorks MATLAB-Simulink Challenge. Prepared by Rohith Ram V, 28 October 2026. This document replaces the pre-review version (Project_Documentation_1.pdf and the v1.9 draft). The README is the source of truth for every number; this document quotes it.

# 1. Summary

The project is an end-to-end breast ultrasound pipeline. Raw OASBUD RF data goes in, a benign or malignant decision comes out, and the pipeline is deployed with NVIDIA Holoscan SDK on an RTX 4070 Laptop GPU. Reconstruction runs as a GPU Coder-generated CUDA library. A second Holoscan application runs plane-wave delay-and-sum beamforming on real per-element channel data (CIRS040GSE) and has no classification stage, because that phantom carries no disease labels.

The MathWorks review found that several headline numbers were measured on data that overlapped training, and that some results could not be reproduced from the committed scripts. This version withdraws those figures and reports held-out results only. The current numbers are: classifier accuracy 81.76% on 636 held-out images (malignant AUC-ROC 0.9105), pipeline accuracy 55.0% (33 of 60) on held-out OASBUD patients, and a CUDA beamformer that matches the scipy reference to 7.64e-13.

# 2. Changes since the review

Each review item was addressed in a commit on the review-response branch. RESPONSE_TO_REVIEW.md maps each item to its commit and files. The main changes are as follows.

| Review finding | Change | Result |
| --- | --- | --- |
| CNN test partition not recoverable; reported figures overlapped training | One dataset builder shared by training and evaluation; split saved with the weights by patient; evaluation loads the saved split | 81.76% held-out accuracy (was 87.37%, withdrawn) |
| 73.0% pipeline accuracy included training patients | Pipeline scored only on the 60 held-out patients | 55.0% (33 of 60) |
| Enhancement filter chosen by PSNR/SSIM against the noisy image | Filters ranked by CNR, gCNR, speckle SNR, edge ratio and a phantom with known ground truth | Bilateral selected (was median) |
| Preprocessing differed between MATLAB and Holoscan | One path: uint8 scaling with no min-max renormalisation | Consistent in both |
| Simulink model described as illustrative | Functional model with a driver script that checks the output | Output 1824 x 510, range 9.55 to 244.25 |
| CUDA source for deployed kernels missing from the repository | Generated CUDA sources committed; personal path fallbacks replaced with errors naming the environment variable | Build instructions match the repository |
| Phased Array Toolbox used only for a 1-D beam | phased.TimeDelayBeamformer swept over 201 angles on real channel data, shown beside the hand-written DAS | 2-D toolbox image in results/figures/Phase5/ |

# 3. Dataset and classifier

The classifier is a MobileNetV2 transfer-learning model. It is trained on a combined dataset of 4,437 images from five sources, with a stratified split of 3,080 training, 721 validation and 636 test images (seed 42). The test partition was saved with the weights, so the evaluation uses exactly the images that were held out from training.

| Source | Images | Benign | Malignant | Normal |
| --- | --- | --- | --- | --- |
| BUSI | 780 | 437 | 210 | 133 |
| BUS-UCLM | 646 | 157 | 71 | 418 |
| BUS-BRA | 1,875 | 1,268 | 607 | 0 |
| BrEaST | 256 | 154 | 98 | 4 |
| OASBUD (PNG export) | 880 | 440 | 440 | 0 |

## Held-out results (636 test images)

| Class | Precision | Recall | F1 | Support |
| --- | --- | --- | --- | --- |
| Benign | 84.72% | 84.27% | 0.8449 | 375 |
| Malignant | 74.65% | 75.00% | 0.7482 | 212 |
| Normal | 90.00% | 91.84% | 0.9091 | 49 |

Overall accuracy is 81.76% (520 of 636). Malignant vs rest: AUC-ROC 0.9105, AUC-PR 0.8154. Single-image MATLAB inference latency is recorded in results/cnn_eval.md and varies between runs.

# 4. Pipeline 1: OASBUD classification

The pipeline runs DataSourceOp, then BeamformingOp (per-column Hilbert envelope, gamma 0.3 power-law compression, GPU Coder CUDA through ctypes), then EnhancementOp (uint8 scaling, bilateral filter, resize to 224 x 224, RGB), then InferenceOp (ONNX MobileNetV2 on CUDAExecutionProvider), then OutputOp. It is run under the GreedyScheduler, so stages execute sequentially.

| Result (60 held-out patients) | Value |
| --- | --- |
| Pipeline accuracy | 55.0% (33 of 60), identical on CUDA and scipy backends |
| Malignant recall | 84.4% (27 of 32) |
| Benign recall | 21.4% (6 of 28) |
| Predictions | 49 of 60 called malignant |
| CUDA vs scipy beamforming | Largest difference 7.64e-13 across 100 patients |
| Throughput, CUDA (committed log) | 35.6 fps mean, 21.8 fps 95th percentile |
| Throughput, scipy (committed log) | 26.9 fps mean |

The low benign recall is most likely a domain gap. The classifier was trained on clinical scanner images, while OASBUD is a single plane-wave acquisition with arc artefacts. The model therefore defaults toward malignant on unfamiliar input. Closing this gap needs training data that resembles OASBUD, and it is listed as an open item.

# 5. Pipeline 2: plane-wave beamforming

HoloscanPlaneWave runs delay-and-sum on genuine per-element channel data from the CIRS040GSE phantom, using a GPU Coder kernel. It has no classification stage, because the phantom has no benign or malignant labels, so none of its results support a diagnostic claim.

| Measure | Committed log | Re-run, 28 Oct 2026 |
| --- | --- | --- |
| CUDA vs numpy reference | 0.00e+00 (5 frames) | 0.00e+00 |
| Data load | 387.74 ms (93% of frame) | 478.88 ms (93%) |
| Beamforming | 8.44 ms (2%) | 8.34 ms (1.6%) |
| Total frame | 461.27 ms | 515.37 ms |
| Throughput | 2.2 fps | 1.9 fps |

The CUDA speedup on beamforming is real, but data loading dominates the frame. Optimising that path is an identified next step and has not been attempted.

# 6. Enhancement filter selection

The earlier comparison scored filters with PSNR and SSIM against the unfiltered noisy image, which rewards doing nothing, so it could not rank denoisers. Filters are now ranked on real OASBUD data (hold-out patient 43mw, fixed lesion and background regions) and on a synthetic phantom with known regions and imnoise speckle at variance 0.05.

| Filter | Real CNR | Real gCNR | Speckle SNR | Phantom PSNR (dB) | Phantom SSIM | Mean rank |
| --- | --- | --- | --- | --- | --- | --- |
| None | 2.0463 | 0.8969 | 4.5276 | 19.18 | 0.1118 | - |
| Median 3x3 | 2.2169 | 0.9234 | 4.8795 | 24.60 | 0.3190 | 2.71 |
| Wiener 5x5 | 2.3059 | 0.9334 | 5.0515 | 29.97 | 0.7120 | 2.14 |
| Bilateral (selected) | 2.5376 | 0.9579 | 5.4701 | 31.66 | 0.7470 | 1.14 |

Bilateral is deployed in both Holoscan pipelines, in Section 9 of the MATLAB evaluation and in the Simulink model. Median keeps the most edge gradient, and at the highest phantom noise level (variance 0.10) Wiener scores above bilateral. Both results are reported in results/enhancement_eval.md.

# 7. Simulink model

Simulink/UltrasoundEnhancementSubsystem.slx is a functional model. A From Workspace block feeds an OASBUD RF matrix through an EnvelopeDetection block (Hilbert envelope, 60 dB log compression, normalised to 0 to 1) and an Enhancement block (the bilateral filter, scaled to 0 to 255), into a To Workspace block. Run it with Simulink/run_enhancement_subsystem.m, which builds the input as [nSamples nElements nTime], calls sim, and checks that the output is finite, the expected size and not all zero. On the first bundled sample patient the output is 1824 x 510 with values from 9.55 to 244.25.

# 8. Reproducing the results

The full command list, with environment variables and expected printed output, is in the Reproduce section of the README. In outline: export the OASBUD PNG training set, train and save the split, evaluate the classifier on the held-out partition, run the enhancement comparison, run the Holoscan benchmarks in WSL2, and run the Simulink driver. Accuracy and verification figures are deterministic for a given model and reproduced exactly on the owner's machine on 28 October 2026. Timings depend on the machine and are recorded in results/logs/rerun_2026-10-28/.

# 9. Limitations and open items

No clinical validation has been done. The classifier result comes from one OASBUD hold-out set with a known domain gap, and benign recall on that set is low. Pipeline 2 has no labels. Frame rates and latencies vary between runs. The committed demo video was recorded before the review changes and is being re-recorded. The CNN was not retrained for this version; the committed weights and split are the ones evaluated. A retrain may differ slightly because of GPU nondeterminism.

# 10. Re-run record, 28 October 2026

Every accuracy and verification figure in the README was reproduced on the owner's machine: classifier held-out accuracy and AUC values (from the saved split), OASBUD hold-out pipeline accuracy on both backends, CUDA verification, Pipeline 2 verification, the enhancement ranking, the FK multi-angle CNR improvement (63.4%), the Simulink output range, and the bundled-sample smoke test (50.0%, 5 of 10). Raw outputs are in results/logs/rerun_2026-10-28/.